OCM Pro CF Radar

Flow rate measurement in pipes and any channel profile

in part filled conditions
Detection of water surface velocity by using the radar

Doppler

Flow calculator for detection and storage of

flow rates calculated from flow velocity and

water level

Water surface velocity detection by using an OFR

radar sensor;

non-contact radar measurement based on the

Doppler principle; measurement of the movement

of waves and calculation of the velocity profile below

the water surface

Temperature compensation of all readings

Indication of measurement values on large back-lit

graphic display

Logging function with trend indication covering 90 days

Clear text programming in dialog mode via front-side

alpha-numeric soft keypad

Operating voltage:	100 – 240 V AC

24 V DC (option)

Constructions:	wall mount enclosure

Protection:	IP 65

Display:	back-lit graphic display

128 x 128 pixel

Inputs:	4 x 4-20 mA, galvanically isolated

4 digital inputs (optocoupler)

Outputs:	4 x 4-20 mA, galvanically isolated

5 x relays (SPDT)

Memory capacity:	up to 128 MB on CompactFlash(

Communication:	Ethernet, analog modem or

(wall mount only)	ISDN; GPRS (please specify), 
ModBus TCP, integrated web server

Version: 	OCP-R4W0 (multifunctional version

with connection option for one radar

sensor)

incl. NivuDat readout software and NivuSoft

evaluation software for Windows 2000/XP/7

deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Surface flow velocity sensor for installation above

channels

Detection of surface flow velocities using the

sensor-integrated radar sensor

Measurement range: 0,15 m/s to 20 m/s

Measurement

uncertainty:	± 2 % of measurement value

Distance:	0,3 m to 20 m

Materials:	plastic, stainless steel

Protection:	IP 68

Cable length:	10 m (option: 15/20/30/50/100 m)

Version:	OFR-F00 001

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 Pc.	UP: ..................

TP: ....................

Combi holder for radar and ultrasonic sensor

Materials:	stainless steel

Version:	ZUB0OFRHAL

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 Pc.	UP: ..................

TP: ....................

Overvoltage protection sensor line

Lightning and overvoltage protection for potential-free

data lines

High performance overvoltage limiter with coarse and

fine protection, with minimum inner resistance, specially

adjusted to LUS and POA sensors

Internal self-control through fail-safe diodes

Maintenance-free since electronically highly 
over-dimensioned

Nominal leakage

surge current (8/20 µs):20 kA

Maximum

operating voltages:	12 V DC/24 V DC

Series resistance:	0.3 Ohm

Fastening:	35 mm DIN rail

Enclosure:	Plastic

Version:	1x DataPro 2x1-12V/12V-11µH

1x SonicPro 3x1-24/24V

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Overvoltage protection analog output

Lightning and overvoltage protection for potential-free

data lines

High-performance overvoltage limiter with

coarse and fine protection for symmetric or

galvanically isolated data line inputs and outputs

as well as 2-wire technology

Suitable for transmission frequencies up to 600kHz

and 0.5A operating current

Internal self control using failsafe diodes

which short-out data lines in case of error.

Maintenance-free due to high electronic

overdimensioning

Nominal leakage

surge current (8/20 µs):20 kA

Maximum

operating voltage:	28 V DC / 20 V AC

Protection rating:	36 V wire-ground

Fastening:	35 mm DIN rail

Enclosure:	Plastic

Version:	DataPro 2x1-24/24V

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Overvoltage protection for power supply lines

High-performance overvoltage limiter

for 230V AC power inputs up to 16 A current.

Overvoltage protection with very high leakage capacity.

Varistor self control and independent disconnection

from mains in case of  exceeding leakage currents.

Visual indication of operation and malfunction

Nominal leakage

current wire-earth:	20 KA

Maximum

operating voltage:	300 V AC

Protection rating:	1.1 kV wire-earth

Fastening:	35 mm DIN rail

Enclosure.	Polycarbonate.

Version:	EnerPro 220 Tr/20

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Accessories

Memory card for mobile data transport

Industry standard with extended temperature range,

suitable for OCM Pro, PCM Pro, PCM 4, NivuSonic

and NivuChannel

suitable to store all single velocities,

average velocity, level / height, flow rate and

temperature as well as additional inputs

Capacity:	128 MB

Temperature range:	-40 °C to 85 °C

Type:	CompactFlash( card 128 MB
Deliver

Type:............................

Make:.....................

Qty.:  2 pc.	UP: ..................

TP: ....................

Read-out unit for Compact Flash™ Card,

to be connected to USB interface

Deliver

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Commissioning of the measurement carried out by

system manufacturer.

Check of hydraulic suitability of installation place,

proper installation of sensors sensor and the

transmitter.

Adjustment and inspection of the measurement

according to VDE/VDI 2640.

Issuing of parameter lists including the device

settings and a commissioning protocol

Qty.:  1 Flat fee	UP: xxxxxxxxxx

TP: ....................