NivuMaster 5 Relais

Measurement transmitter for non-contact level measurement

and volume measurement in open flumes,

channels or on any kind of weirs using

ultrasound.

Fixed calculation operations for linearization of

Venturi, Khafagi or rectangular flumes,

half shells, overflow or V-weirs for volume measurement.

Free programmable characteristic with up to

32 breakpoints. tank volume calculation.

averaging and difference calculation with 2 sensor inputs

trend calculation and pump control

Large back-lit LC display for

alpha-numeric indication of units and for

programming in dialog mode via keypad

(not for 19“) or laptop.

Relays with status LED, can be set to red, yellow

or green for operation or alarm conditions.

DATEM echo analysis with automatic suppression

of false echoes and learning function for error signals.

Time-related power failure protocol.

Transmitter with automatic switchover to emergency

power supply in case of mains failure.

Transmitter

accuracy:	0.25 % (related to distance

between sensor and level)

Resolution:	2mm (0.1% related to the measurement range)

Display:	level and difference simultaneously

Power supply:	230V AC / 115V AC / 24V DC

Input:	1 or 2 ultrasonic sensor(s)

additional 2-wire mA-input (option)

Measurement range:	0.125 to 40 m (depending on sensor)

Outputs:	0/4 - 20mA; load 500 Ohm,

galvanically isolated

RS232 interface for programming

and echo analysis

5 potential-free switchover relays,

free programmable in hysteresis

and function:

min./max. alarm;

in-bounds/out-of-bounds alarm;

trend message decreasing/increasing

with programmable limit value;

error and temperature alarm;

counter for relay switching events;

pump control with and without

pump interchange, cyclic or based on

runtime, group or replacement

operation mode,

start delay as well as

stop delay for complete emptying,

variable switching point to avoid

grease rims,

control functions for flushing dump,

aeration and pump start,

operating hours counter for 5 pumps,

Impulse outputs for external counter

drive at volume measurements and for

time or volume-based sampler drive.

Switching capacity:	5A / 250V AC

Protection:	wall mount: IP65

panel mount: IP65

19”: IP20

Dimensions

(wxhxd):	240x185x120mm (wall)

200x112x110mm (panel)

3HE x 10 TE x 190mm (19”)

Models:	NM0- 5101 65TD E00 (wall)

NM0- 5101 PATD E00 (panel)

NM0- 5101 19HD E00 (19“)

NM0- 5111 65TD E00

(wall mount with add. mA input)

NM0- 5111 PATD E00

(panel mount with add. mA input)

NM0- 5111 19HD E00

(19“ with add. mA input)

Deliver, install and program ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Keyboard for operation and parameter setting

with acoustic entry feedback for NivuMaster 19"

Operation:	membrane keypad

Data transmission:	non-contact via infrared

Model:	programming unit Type NM

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ...................

Analog input converter for adjustment of pressure probes

Type NivuBar Plus II and HydroBar G II to

NivuMaster analog input

to snap onto DIN rail

Power supply:	24 V DC

Analog input:	4-20 mA, active

Analog output:	4-20 mA, passive

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ...................

Surge arrestor for intrinsically safe power supply of

pressure probes Type NivuBar Plus II and HydroBar

G II in Ex-areas zone 1/2
(required in connection with analog input converter)

to snap onto DIN rail

Model: 	9002/13-280-093-001 ATEX

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ...................

Analog output overvoltage protection

Lightning and overvoltage protection for potential-free

data lines according to ATV H 261

High-performance overvoltage limiter with

coarse and fine protection for symmetric or

galvanically isolated data line inputs and

outputs as well as 2-wire technology.

Suitable for transmission frequencies up to 600 kHz

and 0.5A operating current.

Internal self-control using failsafe diodes which

short-out data lines in case of error.

Maintenance-free due to high electric

overdimensioning.

Nominal leakage

surge current

(8/20µs):	20kA

Maximum

operating voltage:	28V DC / 20V AC

Protection rating:	36V wire-ground

Installation:	35mm DIN rail

Enclosure:	plastic

Model:	DataPro 2 x 1 - 24/24V Tr

Deliver and install ready to use

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ...................

Overvoltage protection for power lines

High-performance overvoltage limiter

according to ATV H 261 for 230V AC

power inputs with currents up to 16A.

Overvoltage protection with very high leakage capacity.

Varistor self-control and independent disconnection

from mains in case of exceeding leakage currents

Visible indication of operation and error

Nominal leakage

current wire-ground:	20KA

Maximum

operating voltage:	300V AC

Protection rating:	1.1kV wire-ground

Installation:	35mm DIN rail

Enclosure:	Polycarbonate

Model:	EnerPro 220 Tr/20

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ...................