The tender texts below comprise the most common solutions in practice so that options and parts that are not required can be deleted and only quantities need to be inserted.

For certain projects it is recommended to edit the texts. Our experienced colleagues will be pleased to support you competently, quick and free of charge when it comes to create your tender texts. Please contact us via sales@nivus.com or by phone: +49 7262 9191-794.

NivuFlow 7550

Hybrid, Ex Zone 1-protected flow measurement of

slight to heavily polluted media in various full and

temporarily part filled channel profiles and pipes.

Redundant combined velocity measurement method 
for contactless measurement of the surface velocity 
and simultaneous medium-contacting measurement

of real flow profiles via spatially allocated individual

flow velocities using the cross correlation method

covering the entire wetted cross-sectional area.

False readings and measurement failures caused by 
backwater, extremely low flow velocities down to

standstill and immersed radar sensor are considered

and eliminated if required.

Reliable flow velocity measurement even at very high
velocities, shooting discharge and lowest flow levels 
of only a few millimetres.

Hybrid flow transmitter with integrated 
numeric discharge models for determination of average

flow velocities from the surface velocity and up to

16 spatially allocated individual flow velocities

within the medium using the cross correlation method.

A surface radar sensor and a medium-contacting

flow velocity sensor to determine the accurate flow

profile distribution in various part filled and full

channel profiles and pipes can be connected

simultaneously.
Detection of spatially allocated flow velocity profiles

using the correlation method and digital pattern detection.
Calculation of the real 3D flow profile through
integration of all measured surface velocities and spot

velocities covering the wetted cross section.

Calculation of the flow rate from  average flow velocity,

filling level and channel shape considering medium

temperature, channel shape, channel dimensions and

level of partial filling based on internal numeric discharge

models. Storage, graphic representation and transmission

ofall relevant measurement data.

Determination of water surface level using ultrasound
from bottom up, sensor-integrated pressure cell, ultrasound

from top down or via multiple separate level measurements.

Various level measurement methods can be combined

for optimum measurement adaptability to high

measurement dynamics and difficult hydraulic

applications and for redundant recording.

Energy-saving cycle operation mode for DC measurement systems

without external power supply with free selectable 
measurement/sleep cycles.
Switchover between cycle mode to event mode

for high-resolution detection of critical flow rates digitally

controlled and based on limit values.
Event mode free programmable for short cycle or permanent
measurements.

Metal-encapsulated compact enclosure with easily accessible

plug-in spring cage terminals on the front panel for

quick installation and mounting.

Large daylight TFT full colour graphic display featuring

240 x 320 px and 65536 colours.

Optimum display clarity even in direct sunlight thanks to

1:400 contrast ratio and up to 80° possible deviation

from observation plane in any direction.

Modern and intuitive one-hand operation using rotary

pushbutton in context with high-resolution graphic

colour display.

Programming with flat hierarchical structure, supported

by colours and graphic elements, visual true-to-scale

representation even of complex canal shapes and sensor 
positions in order to avoid errors during commissioning.

Intelligent error and sensor diagnostics.

Integrated recorder function with min. and max. memory.

Storage of all individual velocities and associated

position levels, flow levels, average flow velocities,

flow rates, medium temperature and air temperature in

large internal memory.

Programming data and all saved readings as well as errors

can be copied to standard USB stick via front-side

USB interface.

HART interface for direct connection of compatible

sensors.

Connection of OCL sensors and v-sensors via

RS232/485 interface.

Connection of the transmitter to following systems using

HART, RS232/485, TCP-IP/Ethernet and ModBus TCP/RTU.

Worldwide connectivity via mobile radio through

integrated MultiRoam SIM chip as well as additional

SIM card slot (option).

Direct transmission of measurement and system data,

error messages and alarms via mobile radio to the NIVUS

web portal.

Remote maintenance, remote diagnostics and remote

commissioning with the full scope of services of the on-site

commissioning.

Operating voltage: 	85 – 240 V AC

9 – 36 V DC (optional)

9 – 36 V DC cycle operation
	(optional)

Construction:	DIN rail

Enclosure material:	profile aluminium

Protection:	IP 20

Ex Protection:	Zone 1

(optional)	using Ex separator Type iXT

Display:	full colour graphic TFT daylight display

240 x 320 px, 65536 colours

contrast min. 1:400

Programming:	dialog mode using rotary pushbutton

and 2 function keys

Version: 	N75-5KR1

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Licence Key

For device-related transmission of measurement data 
via FTP and/or e-mail to customer server or free
selectable e-mail address.

Transfer of the serial number-based access data for

the activation of a NivuFlow measurement

transmitter by the operator.

Version:	NFX0 LIZENZ FTP

Transmit access data via e-mail

Type: ............................

Make: .....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Robust, corrosion-proof field enclosure with large

transparent cover for NF 7550 and one optional

Ex-Separation Module.

Special enclosure with enhanced protection degree

for outdoor installation of the transmitter.

Material:	plastic

Protection:	IP 67 or
	IP 68 (optional)

Versions:	ZUB0 NFW0

ZUB0 NFW0 IP68

Deliver incl. accompanying cable glands, install

transmitter and isolating amplifier if required and

install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Module for safe separation of Ex areas between

Ex-compatible sensors and Type NF7550 transmitters.

Suitable for direct connection of Type POA, CS2 and

OFR flow velocity sensors; Electronic Box Type EBM,

Type LUS level sensors and NivuBar Plus.

Ex-protected HART-interface for direct connection

of i-Series level sensors

Ex module can be installed in a distance of up to 300 m

from the transmitter in separate protective enclosure.

Operating voltage: 	12 V DC (from NivuFlow 7550)

Construction:	DIN rail

Enclosure material:	plastic

Protection:	IP 20
Ex Protection:	Zone 1

Max. distance

Sensor – iXT:	150 m

Max. distance

iXT – NF750:	300 m

Versions: 	iXT0 420 (2x v, 1x LUS, 1x mA)

iXT0 421 (2x v, 1x LUS,
1x HART)

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Flow-optimised flow velocity sensor for installation on

channel bottom, channel side or floating installation

from top down.

Detection of up to 16 spatially allocated flow velocities

using built-in ultrasonic sensor.

Measurement range:	- 100 cm/s to 600 cm/s

Meas. uncertainty:	1 % of measurement range

Materials:	stainless steel 1.4571; PPO GF30

PEEK, PE, PU

Protection:	IP 68

Ex Approval:	Zone 1 (option)

Cable length:	10 m (option: 15/20/30/50/100 m)

extendable up to 150 m

Version:	POA-V200KT

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Flow-optimised flow velocity sensor with built-in

pressure measurement cell for installation on

channel bottom or channel side.

Detection of up to 16 spatially allocated flow velocities

using built-in ultrasonic sensor.

Additional detection of flow levels using built-in

pressure measurement cell with high chemical resistance

Measurement ranges

Flow velocity:	- 100 cm/s to 600 cm/s

Pressure cell:	0 – 350 cm  level

Meas. uncertainties

Flow velocity:	1 % of measurement range

Pressure cell:	max. 0.75 % zero point drift of MR

max. 0.5 % of final value

Materials:	stainless steel 1.4571; PPO GF30

PEEK, PE, PU

Protection:	IP 68

Ex Approval:	Zone 1 (option)

Cable length:	10 m (option: 15/20/30 m)

extendable up to 150 m using

pressure compensation element

Version:	POA-V2D0K

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Flow-optimised combi sensor with built-in

pressure measurement cell for installation on

channel bottom.

Detection of flow level and up to 16 spatially allocated

flow velocities using two built-in independent ultrasonic

sensors.

Additional detection of flow levels using built-in

pressure measurement cell with high chemical resistance

Measurement ranges

Flow velocity:	- 100 cm/s to 600 cm/s

Water ultrasound	4 – 200 cm

Pressure cell:	0 – 350 cm

Meas. uncertainties

Flow velocity:	1 % of measurement range

Water ultrasound:	better than 2 mm

Pressure cell:	max. 0.75 % zero point drift of MR

max. 0.5 % of final value

Materials:	stainless steel 1.4571; PPO GF30

PEEK, PE, PU

Protection:	IP 68

Ex Approval:	Zone 1 (option)

Cable length:	10 m (option: 15/20/30 m)

extendable up to 150 m using

pressure compensation element

Version:	POA-V2U1K

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Flow velocity sensors for average flow levels

higher than 1 m

Flow-optimised flow velocity sensor for

installation on the channel bottom, channel side or for

floating installation

Detection of up to 16 individual, spatially allocated flow

velocities using built-in ultrasonic sensors

Measurement range:	- 100 cm/s to 600 cm/s

Meas. uncertainty:	1 % of measurement range

Materials:	stainless steel 1.4571; PPO GF30

PA6 GF30, CFK, PU

Protection:	IP 68

Ex approval:	Zone 1 (option)

Cable length:	10 m (option: 15/20/30/50/100 m)

extendable up to 250 m

Model:	CS2-V200 KT

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Flow-optimised flow velocity sensor for

installation on the channel bottom or channel side

Detection of up to 16 individual, spatially allocated flow

velocities using built-in ultrasonic sensor

Flow level detection using sensor-integrated

pressure measurement cell with high chemical resistance

Measurement ranges

Flow velocity:	- 100 cm/s to 600 cm/s

Pressure cell:	0 – 350 cm

Meas. uncertainties

flow velocity:	1 % of measurement range

Pressure cell:	max. 0.75 % zero point drift/MR

max. 0.5 % of final value

Materials:	stainless steel 1.4571; PPO GF30

PA6 GF30, CFK, PU

Protection:	IP 68

Ex approval:	Zone 1 (option)

Cable length:	10 m (option: 15/20/30 m)

via pressure compensation element 
extendable to up to 250 m

Model:	CS2-V2D0 KT

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Flow-optimised flow velocity sensor with additional 
pressure measurement cell for installation on the channel

bottom

Detection of up to 16 individual, spatially allocated flow

velocities using built-in ultrasonic  sensor

Flow level detection using two redundant ultrasonic

sensors as well as an additional sensor-integrated

pressure measurement cell with high chemical resistance

Measurement ranges

Flow velocity:	- 100 cm/s to 600 cm/s

Water ultrasound:	12 – 400 cm

Pressure cell:	0 – 350 cm

Meas. uncertainties

flow velocity:	1 % of measurement range

Water-ultrasound:	better than 2 mm

Pressure cell:	max. 0.75 % zero point drift/MR

max. 0.5 % of final value

Materials:	stainless steel 1.4571; PPO GF30

PA6 GF30, CFK, PU

Protection:	IP 68

Ex approval:	Zone 1 (option)

Cable length:	10 m (Option: 15/20/30 m)

via pressure compensation element 
extendable to up to 250 m

Model:	CS2-V2U1 KT

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Radar sensor

Surface flow velocity sensor OFR for

installation above channels

Detection of the surface flow velocity using

the sensor-integrated radar sensor

Measurement range:	0.15 m/s to 10 m/s

Material:	plastic

Protection:	IP 68

Cable length:	10 m (optional: 15/20/30/50/100 m)

Versions:	OFR-G000 0101 0

OFR-EV00 0101 0 (Ex version)

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Ultrasonic level sensor with integrated temperature

sensor to compensate ultrasound transit time and

built-in free programmable evaluation electronic.

Encapsulated, permanently immersible sensor in
corrosion-proof enclosure with high chemical

resistance.

Suitable for measurement of liquids and

pasty media.

Measurement range:	0.3 to 6 m

Meas. uncertainty:	0.25%

(distance sensor-surface)

Resolution:	2 mm

Startup time:	4 sec. until output of readings

Memory:	Flash

Calibration:	via HART

Enclosure:	Valox 357; (optional: PVDF)

Temperature range:	-40° to max. +80 °C

Fastening:	rear thread 1“

optional: front thread 1.5“

Protection:	IP 68

Immersion	with SUB sleeve (optional)

detection:

Ex Protection:	Ex mb for Zone 1

Ex ia for Zone 0 (optional)

Output:	4 - 20 mA / HART

Cable length:	10 m (optional: 20/30 m)

extendable up to 1000 m

Version:	NMI0 06

Deliver, install ready to use and calibrate

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Sensor line overvoltage protection

Lightning and overvoltage protection for potential-free

data lines, according to DWA A-M 261 suitable as

overvoltage protection for wastewater treatment plants

High-performance overvoltage limiter with

coarse and fine protection, with minimum internal

resistance, particularly adjusted to LUS and POA sensors

Internal self-control using failsafe diodes

Maintenance-free due to high electronic

Nominal leakage

surge current

(8/20 µs):	20 kA

Max. operation

voltage:	12 V DC/24 V DC

Series resistance:	0.3 Ohm

Installation:	35 mm DIN rail

Enclosure:	plastic

Ex approval:	Zone 1 (option)

Models:	1x DataPro 2x1-12V/12V

1x SonicPro 3x1-24/24V

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Satz.	EP: ..................

GP: ....................

Analog output overvoltage protection

Lightning and overvoltage protection for potential-free

data lines

High-performance overvoltage limiter with

coarse and fine protection for symmetric or

galvanically isolated data line inputs and

outputs as well as for 2-wire systems

Suitable for transmission frequencies up to 600 kHz

and 0.5 A operating current

according to DWA A-M 261 suitable as

overvoltage protection for wastewater treatment plants

Internal self-control using failsafe diodes

which short-out data lines in case of error.

Maintenance-free due to high electronic

overdimensioning

Nominal leakage

surge current

(8/20 µs):	20 kA

Maximum

operating voltage:	28 V DC / 20 V AC

Protection rating:	36 V wire-ground

Installation:	35mm DIN rail

Enclosure:	plastic

Ex approval:	Zone 1 (option)

Model:	DataPro 2x1-24/24V

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Power lines overvoltage protection

High-performance overvoltage limiter

for 230 V AC power inputs with 16 A current,

according to DWA A-M 261 suitable as

overvoltage protection for wastewater treatment plants.

Overvoltage protection with very high leakage capacity.

Varistor self-control and independent disconnection

from mains in case of  exceeding leakage currents.

Visual indication of operation and malfunction

Nominal leakage

current wire-ground:	20 KA

Maximum

operating voltage:	300 V AC

Protection rating:	1.1 kV wire-ground

Installation:	35 mm DIN rail

Enclosure:	Polycarbonate

Model:	EnerPro 220 Tr/20

Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Accessories

Pressure compensation element

For connection or to extend sensors

Types V2U and V2D.

Consisting of junction box made of metal incl. clamps,

Pressure compensation plug and socket as well as large

air filter with visual self-control for absorption

of air humidity.

Enclosure:	Al die-cast

Protection:	IP 65 (enclosure)

IP 40 (filter)

Type:	ZUB0 DAE

Deliver and install ready to use

Type:............................

Make:.....................

Qty.:  1 pc.	UP: ..................

TP: ....................

Corrosion-proof combi holder bracket for radar sensor 
and ultrasonic sensor

Material:	1.4301

Versions:	ORH-A00

(modular holder bracket for radar
sensor)
ORH-AL0
(modular holder bracket for radar
sensor and i-Sensor)
ORH-ALV6
(modular holder bracket with cover for radar
sensor and i-Sensor)

Incl. corrosion-proof incidentals and fastening material
Deliver and install ready to use

Type:............................

Make:.....................

Quantity:  1 Pc.	UP: ..................

TP: ....................

Initial start-up and function al check of measurement 
system carried out by manufacturer.

Verification of hydraulic suitability of the place of

installation, correct sensor mounting and installation

of transmitter.

Handover of parameter lists containing implemented

instrument setup and initial start-up protocol.

Prerequisites for commissioning:

Sensors installed and transmitter(s) connected ready to

use according to manufacturer specifications

Quantity:  1 flat rate	UP: xxxxxxxxxx

TP: ....................